# Checklist for Preclinical Studies

# Medical Device

Notes

This workbook separates the medical device checklist into category-specific tabs: Preclinical studies confirm biological safety, mechanical, chemical, and functional properties of a device before use in humans. Note: Completion of all studies is not required; please indicate which studies have been completed.

References include ISO 10993 series, IEC 60601 series, ISO 14971, IEC 62304, IEC 62366, ISO 11135 / 11137 / 17665 / 11607, ASTM Standards, FDA 21 CFR Parts 807, 812, 820, 21 CFR Part 58, FDA guidance documents.

| Category | Item/Study | Purpose | Key Outputs/Acceptance Criteria | Reference/Guideline | Notes |
| --- | --- | --- | --- | --- | --- |
| Biocompatibility Assessment | ☐Cytotoxicity | Forms the basis for biological risk assessment |  | ISO 10993-5 | Scope of testing must be justified based on risk analysis and, where possible, supported by literature data and material compliance information. |
| Biocompatibility Assessment | ☐Irritation and sensitization | Depending on the expected contact with the body, anatomical location, exposure time, and materials used. |  | ISO 10993-10 | Same as above |
| Biocompatibility Assessment | ☐Local and systemic toxicity | Same as above |  | ISO 10993-11 | Same as above |
| Biocompatibility Assessment | ☐Mutagenicity and carcinogenicity | Same as above |  | ISO 10993-3 | Same as above |
| Mechanical Testing | ☐Strength ☐Resistance to damage and fatigue | Verify the device’s resistance to forces that may occur during use. For implants, joint replacements, or surgical tools, testing may include fatigue strength, fracture resistance, stiffness, and deformation under force. |  | FDA Guidance: Recommended Content and Format of Non-Clinical Bench Performance Testing Information in Premarket Submissions (2019) | Bench Performance Testing |
| Functional Testing | ☐Flow ☐Blockage ☐Leakage ☐Compatibility with other systems/devices | Simulating clinical conditions in the lab environment |  |  | Bench Performance Testing |
| Chemical and Toxicology testing | ☐Chemical Composition ☐Substance mitigation ☐Processing Residues ☐Corrosion Testing ☐Fatigue Testing ☐Wear Testing | Chemical analysis of materials to identify components that may affect biocompatibility, performance, and safety standards. Extensive extractables/leachables analysis Material degradation Fatigue, debris generation |  |  | Common for implants or moving components |
| Process Validation | ☐Sterilization Validation | Full Validation |  | ISO 11135 (ETO) ISO 17665 (steam) or ISO 11137 (radiation) | If sterile |
| Process Validation | ☐Package Integrity Validation ☐Shelf-Life Validation | Complete sterile barrier validation Real-time and accelerated aging |  | ISO 11607 ASTM F88 ASTM F1980 |  |
| Process Validation | ☐Chemical Residues ☐Endotoxins ☐Processing Agents | Confirm effectiveness of processes |  |  |  |
| EMC Testing | ☐Radiated & conducted emissions ☐Radiated & conducted immunity ☐ESD, EFT, and surge ☐Voltage dips & interruptions ☐Magnetic field immunity | Medical devices must perform safely in real-world electromagnetic environments |  | IEC 60601-1-2 | If powered |
| Human Factors and Usability Testing |  | Ensure that a medical device is effective, safe, and user-friendly. Formal Validation study | FDA Guidance: Applying Human Factors and Usability Engineering to Medical Devices (2016, current) | IEC 62366 |  |
| Cybersecurity and Software Device Requirements | ☐Data breach ☐Unauthorized access ☐Software Verification/validation | Comprehensive documentation Full security risk management | FDA Guidance: Cybersecurity in Medical Devices: Quality System Considerations and Content of Premarket Submissions (2023) FDA Guidance: Content of Premarket Submissions for Device Software Functions | IEC 62304:2006 | If software present |
| Environmental Testing | ☐Humidity ☒Pressure ☒Shock ☐Temperature ☐Vibration |  |  |  |  |
| Integration Testing |  | Verifying the interactions and interoperability between a medical device and other components or systems it communicates with. |  |  |  |
| Animal Testing | ☐Acute studies ☐Chronic studies ☐Histology ☐Local tissue response ☐Functional performance | GLP Animal Studies |  | FDA Guidance: General Considerations for Animal Studies Intended to Evaluate Medical Devices (2023) | Frequently required but may be substituted for phantom testing or equivalent bench top testing with FDA consensus. |